## Sheet1
FMEA
Various parts of the system can fail. For each part, this analysis lists the consequences.
Problem | Consequence | Motor alarm (Sparta0 model) | Tidal alarm (mkII) | Low PEEP alarm (mkII) | High PEEP alarm (mkII) | Low PP alarm (mkII) | High PP alarm (mkII) | Severity | Likeliness | Detection % (mkII) | Risk prio w/o detection | Spartan model Risk priority (RPN) | MkII model Risk priority (RPN) | Comment
Disconnected patient | No air to patient | 0 | 0 | 100 | 0 | 100 | 0 | 100 | 80 | 100 | 80 | 80 | 0
Blocked airway in INSP connection to patient | No inspiration | 0 | 10 | 10 | 0 | 100 | 0 | 100 | 50 | 100 | 50 | 50 | 0
Blocked airway in EXP connection from patient | No expiration | 0 | 0 | 0 | 100 | 100 | 100 | 100 | 50 | 100 | 50 | 50 | 0
Pressure sensor hose at patient disconnected | No pressure data | 0 | 0 | 100 | 0 | 100 | 100 | 50 | 50 | 100 | 25 | 25 | 0
Inlet blocked | Low pressure buildup | 5 | 0 | 0 | 0 | 100 | 0 | 100 | 5 | 100 | 5 | 4.75 | 0
Exhaust blocked | Expiration insufficient | 0 | 0 | 0 | 100 | 100 | 0 | 100 | 5 | 100 | 5 | 5 | 0
Pump leaky | Lower pressure, air not filtered | 0 | 0 | 0 | 0 | 100 | 0 | 50 | 20 | 100 | 10 | 10 | 0
Motor fails due to wear | No pressure | 100 | 100 | 50 | 0 | 100 | 0 | 100 | 50 | 100 | 50 | 0 | 0
Linkages break due to wear | No pressure | 0 | 100 | 50 | 0 | 100 | 0 | 100 | 50 | 100 | 50 | 50 | 0
Electronics stops, not powering the motor | No pressure | 100 | 100 | 50 | 0 | 100 | 0 | 100 | 20 | 100 | 20 | 0 | 0
Electronics fails, keeping power on | Too high rate | 0 | 100 | 0 | 0 | 0 | 10 | 100 | 5 | 100 | 5 | 5 | 0
One-way valve from SUPPLY to PUMP port leaks | Lower pressure | 0 | 0 | 0 | 0 | 100 | 0 | 20 | 20 | 100 | 4 | 4 | 0
One-way valve from SUPPLY to PUMP port stuck open | No pressure | 0 | 100 | 50 | 0 | 100 | 0 | 100 | 2 | 100 | 2 | 2 | 0
One-way valve from SUPPLY to PUMP port stuck closed | No pressure | 0 | 100 | 100 | 0 | 100 | 0 | 100 | 2 | 100 | 2 | 2 | 0
One-way valve from PUMP to INSP port leaks | Pressure drops | 0 | 0 | 50 | 0 | 100 | 0 | 20 | 20 | 100 | 4 | 4 | 0
One-way valve from PUMP to INSP port stuck open | Air sucked back into pump, CO2 rises | 0 | 0 | 50 | 0 | 20 | 0 | 100 | 2 | 50 | 2 | 2 | 1
One-way valve from PUMP to INSP port stuck closed | No pressure | 0 | 100 | 100 | 0 | 100 | 0 | 100 | 2 | 100 | 2 | 2 | 0
One-way valve from SUPPLY to INSP port leaks (or both valves near PUMP leak) | Lower pressure | 0 | 0 | 50 | 0 | 100 | 0 | 50 | 20 | 100 | 10 | 10 | 0
One-way valve from SUPPLY to INSP port stuck open | No pressure | 0 | 50 | 50 | 0 | 100 | 0 | 100 | 2 | 100 | 2 | 2 | 0
One-way valve from SUPPLY to INSP port stuck closed | Autonomous breathing heavier | 0 | 0 | 0 | 0 | 0 | 0 | 10 | 20 | 0 | 2 | 2 | 2
One-way valve from PUMP to OVP port leaks (not present in Spartan Acrylic model) | More water level changes | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 20 | 0 | 0 | 0 | 0
One-way valve from PUMP to OVP port stuck open (not present in Spartan Acrylic model) | More water level changes | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0
One-way valve from PUMP to OVP port stuck closed (not present in Spartan Acrylic model) | Pressure may become too high | 0 | 0 | 0 | 10 | 0 | 100 | 50 | 20 | 100 | 10 | 10 | 0 | An other valve, from EXP to OVP, will still protect
One-way valve from EXP to OVP port leaks | More water level changes | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 20 | 0 | 0 | 0 | 0
One-way valve from EXP to OVP port stuck open | More water level changes | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0
One-way valve from EXP to OVP port stuck closed | Pressure may become too high, coughing blocked | 0 | 0 | 0 | 80 | 0 | 100 | 80 | 20 | 100 | 16 | 16 | 0
One-way valve from EXP to PEEP port leaks | More water level changes | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 20 | 0 | 0 | 0 | 0
One-way valve from EXP to PEEP port stuck open | More water level changes | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0
One-way valve from EXP to PEEP port stuck closed | No expiration, pressure rises | 0 | 10 | 0 | 100 | 0 | 0 | 100 | 20 | 100 | 20 | 20 | 0
Transistor valve from EXP to PEEP port leaks | Lower PIP | 0 | 10 | 0 | 0 | 100 | 0 | 50 | 20 | 100 | 10 | 10 | 0
Transistor valve stuck open | PEEP stays, no inspiration | 0 | 20 | 0 | 0 | 100 | 0 | 100 | 20 | 100 | 20 | 20 | 0
Transistor valve stuck closed | No expiration, pressure rises | 0 | 20 | 0 | 100 | 0 | 0 | 100 | 20 | 100 | 20 | 20 | 0
Transistor valve membrane broken | No pressure buildup | 0 | 50 | 10 | 0 | 100 | 0 | 100 | 20 | 100 | 20 | 20 | 0
Water level too low | Pressures too low | 0 | 0 | 100 | 0 | 50 | 0 | 50 | 50 | 100 | 25 | 25 | 0
Water level too high | Pressures too high | 0 | 0 | 0 | 100 | 0 | 50 | 50 | 50 | 100 | 25 | 25 | 0